## Slide 1
6/15/2018
1
EPICS Collaboration Mtg.
APS, Argonne/IL June. 15, 2018
On-axis 3D Microscope for X-ray Beamlines at NSLS-II
Kazimierz J. Gofron (kgofron@bnl.gov), NSLS-II

## Slide 2: Outline
Overview –NSLS2
  Beamlines (26 operational + 3 under constr.)
  End stations
  Beam size(s) @ sample 1 [um], 4 [um], 4x6[um]
  Sample/features: 1 [um]
On-axis microscope(s): Non-dispersive (mirror) optics
Computer Vision
  Goniostat centering (17-ID, 19-ID)
  X-ray beam detection, stability studies
  Real time image processing
  Focus stacking
  3D microscope
  AD barcode plugin
Credits

## Slide 3
X-ray
M
Objective
Cam
Sample
K.J. Gofron, et. al., NIMA, Vol. 649, Issue 1,
1 September 2011, Pages 109-111.
On-axis beamline microscopes
IXS (10ID)
WD=60 cm
3 [um]
3 x
AMX/FMX
(17ID)
30x,5x
1 [um]
Diamond cell (IXS)
X-ray

## Slide 4
Endstation microscopes (IXS)

## Slide 5: Endstation microscopes (TXM)
2x
10x
5x
1x
All are home assembled microscopes
2 X-ray detecting microscopes
2 optical alignment microscopes

## Slide 6
Endstation microscopes (FXI)
Sample imaging
X-ray imaging , TXM

## Slide 7
X-ray eye: 10ID (KB mirror - VFM)
KB mirror (VFM)
Cam
X-ray
20x
Focus
Mirror
CWO
7 um
FWHM = 5 um

## Slide 8: Computer Vision
Computer Vision for beamline with openCV:
Development of image analysis software backed by OpenCV
Development of easy use python module to access OpenCV functions
Optimized results for fast computation via C/C++ backed code, along with Intel IPP/TBB libraries.
GPU role.
Camera -> areaDetector { | IP based} -> Comp. Vision -> Results {sorted size, intensity, position, pattern,…} -> EPICS PV { | file | control software | …}
Automate processes such as:
  Positon, spread, and intensity of X-Ray Beams
  Isolate crystals and X-Ray streaks
  Provide assistance to sample mounting
  Calibrate Goniostat Rotation and Robotic Vision
3D microscope

## Slide 9: Computer Vision – loop center (5s)
[um]

## Slide 10: Computer Vision
Console Output:
Object Details:
perimeter: 2356.99022925
orientation: 179.838363647
max: (925, 198)
height: 372
extrema: {'B': (938, 568),
'R': (1054, 415),
'L': (813, 377),
'T': (914, 196)}
area: 65058.5
min: (1047, 564)
sum intensity: 20426526
width: 241
centroid: (933, 382)
mean intensity: 227.842390577
IXS – BPM1
IXS Point & Click: Left: Image Result with Contour and Top Extrema
Right: Point & Click GUI Interface ( X, Y, MC Scale Bars, Green Cursor)
Console Output: PIN TOP: (1723, 1306)
Console Output:
Object 1:
perimeter: 125.840619564
orientation: 179.981033325
max: (131, 78)
height: 55
extrema: {'B': (129, 122), 'R': (135, 98),
'L': (126, 92), 'T': (132, 67)}
area: 270.5
min: (134, 83)
sum intensity: 62689
width: 9
centroid: (130, 95)
mean intensity: 126.644444444
IXS: Merlin Data Results for First (Largest) Object
AMX: Robot gripper – sample detection

## Slide 11
Resolution and Depth of Focus
Resolution R = wavelength / (2*NA)
Depth of field = wavelength*n/(NA)^2; n=1 for air
Depth of field = 4*n*R^2/(wavelength)
Shallow DoF is a Weakness
Resolution [um]
DoF [um]

## Slide 12: Focus – Horse Fly
?

## Slide 13: How it Works
To begin, a set of images is taken by a shallow depth of field camera, all of the same sample.
The images must then be aligned, in order for the merge to create a seamless image.
A Euclidean transform was the alignment method used, after Oriented-Fast and Rotated-Brief proved ineffective when working with blurry images.

## Slide 14: How it Works
Next, each image is blurred with a Gaussian blurring filter, to further emphasize the in-focus regions of the image
Finally, a Laplacian kernel is run across each image, assigning a sharpness value to each pixel.

## Slide 15: How it Works
Finally, we iterate over each pixel choosing the one from the image with the highest “sharpness” assigned to it
Additionally, during this step, we take note of which image provided each pixel, effectively creating a depth map of the image.

## Slide 16: How it Works
Finally, using the depth map, a 3D model is constructed by a custom OpenGL pipeline, and the merged image is placed over this model.
The 3D model, once rendered is rotated, and made into a video for easier evaluation.
Additionally, using the EPICS python wrapper, automated image stack collection was implemented for supported cameras.

## Slide 17: Focus Stacking – Horse Fly (5x)
Original 38 Image Stack
ECC
Gaussian Blur/Laplacian
Alignment
Edge Detection
ORB

## Slide 18: The code
Real Time

## Slide 19: Key Points
Code portability was important: Python chosen as coding language because of this and because of ability to call efficient C and C++ functions through wrappers
OpenCV and OpenGL used for Computer Vision and Graphics respectively due to them being powerful, open source, and compatible with all major operating systems.
Python wrappers for OpenCV and OpenGL used to improve performance over native Python code

## Slide 20: Beettle (10ID, 3x/3um)
Two examples out of a stack of 19 images. First we performed focus stacking and gamma correction. Images were taken at the IXS 10ID beamline.
https://www.youtube.com/channel/UC-SfBpwDIiuw41_r0qqYkZQ?view_as=subscriber

## Slide 21: The Beetle – merged/gamma corrected (3x)
10ID

## Slide 22: The Beetle – depth map

## Slide 23: The Beetle: focus=3rd dimension
https://www.youtube.com/watch?v=UFx2EDouO-k

## Slide 24: Mineral
https://www.youtube.com/watch?v=FdpdAdoirwA

## Slide 25: Crystal harvesting – sonic ejection
Crystal images are all courtesy of Alexei Soares and the Click to Mount Team. Focus stacks of 15 images.
Crystal are sonicly ejected only from top layer, and surface layer moves (with crystals) to “heal” the surface.
Crystals from lower layers are not harvested.

## Slide 26: Crystal harvesting
Depth map
L1 Surface
L2 Middle
~10 um
Ln Deep
………
Sonic harvesting of crystals
is driven by surface tension.
Conclusion 1:
Prevent crystals from settling to the bottom. Thus consider using Bingham fluid solutions.
(to be published)
SUMMARY: Converted weakness (shallow DoF) into strength (recovered 3rd depth dimension)
Crystals are ejected from surface layer
Sonic source

## Slide 27: AD Plugin: Barcode reader
Enabling technology
Open source
- openCV
- zbar

## Slide 28: References
Moeller, Michael, et al. “Variational Depth From Focus Reconstruction.” IEEE Transactions on Image Processing, vol. 24, no. 12, 2015, pp. 5369– 5378.,doi:10.1109/tip.2015.2479469.
“OpenCV Library.” OpenCV Library, opencv.org/.
Group, Khronos. “The Industry's Foundation for High Performance Graphics.” OpenGL.org, www.opengl.org/.

## Slide 29
Work at Brookhaven was supported by the Department of Energy, Office of Basic Energy Sciences under contract DE-AC-02-98CH10886
Credits
Jakub Wlodek (focus stack)
Yong Cai (10ID)
Bill Watson (comp. vision)
Scott Coburn (mechanical)
Stephen Antonnelli (mechanical)
Bruno Martins
Martin Fuchs (17ID)
Jean Jakoncic (17ID)